APPROVED BY

Lenar Valeev

General Director

EVOTECH-MIRAI GENOMICS LLC

________________

__________________ 20____

DATASHEET No.____

of ____________

Isothermal amplifier for 
SARS-CoV-2 nucleic acid detection according to TU 26.51.53-003-06931260-2020

Registration Certificate for a Medical Device

No. 2020/10090 of April 17, 2020

2020

MANUFACTURER

Manufacturer:

EVOTECH-MIRAI GENOMICS Limited Liability Company

(EVOTECH-MIRAI GENOMICS LLC)

Address: 3B Akademika Artsimovicha St, Room 11, Moscow 117437, Russia

The production site is located at:

Joint-Stock Company «Production Association «Zavod imeni Sergo» (POZIS) JSC POZIS, 4, Privokzalnaya Street, Zelenodolsk, Republic of Tatarstan, Russia 422546

LIST OF EQUIPMENT

| Pos. | Item | Quantity |
| --- | --- | --- |
| 1.1 | Isothermal amplifier for SARS-CoV-2 nucleic acid detection according to TU 26.51.53-003-06931260-2020 (the “Amplifier”) | 1 |
| 1.2 | USB 2.0 High-speed А-В cable | 1 |
| 1.3 | Power supply unit | 1 |
| 1.4 | USB drive with software | 1 |
| 1.5 | Operational documents: |  |
| 1.5.1 | Handbook | 1 |
| 1.5.2 | Datasheet | 1 |
| 1.6 | Transport case (plastic protective case) | 1 |
| 2.1 | Accessories: |  |
| 2.1.1 | Laptop PC (optional) | 1 (optional) |

SPECIFICATIONS

Amplifier weight: 2.5±0.1 kg.

The Amplifier operates on AC. Power supply unit specifications: input voltage is АС 100-240V at the frequency of 50/60 Hz, output voltage is DC 12V at 6А

Maximum nominal rated input: 6A.

The Amplifier includes three independent thermal blocks (channels).     Each thermal block is designed to fit 8 standard 0.2ml PCR tubes.

The Amplifier specifications are shown in Table 1.

Table 1.

| Specification | Value |
| --- | --- |
| Thermal block format | 24 x 0.2ml tubes (8 x 3) |
| Reaction mixture volume | 10-50μl |
| Thermal block temperature | 66.0 - 67.0°С |
| Thermal block temperature tolerance | ±0.1°С |
| Heated lid temperature | Lower than 70°С |
| Thermal block heating device | Resistive heater |
| Excitation source | LED |
| Sensor | Photodiode |
| Number of channels for fluorescence measurements | 1 |
| Excitation/detection wavelength | 464 nm / 520 nm |
| Computer interface | USB 2.0 full speed |

The Amplifier is controlled by Amplifier Client V.1.0.0.0 software installed on a PC meeting the following minimal requirements:

- OS: Windows 7, Windows 8, Windows 10

- Pentium IV 1,000 MHz or higher

- minimum 512MB of RAM

- minimum 100MB free disk space

- CD drive

- display resolution of at least 1024x768

- available USB 2.0 High-speed ports.

The maximum allowable time to establish the operating mode of the Amplifier after switching on is no more than 10 minutes.

The maximum nonstop continuous run time is 8 hours, the minimum recovery time is 30 minutes.

Catalogue number

A5506 - Medical Device

А5506RUO - Research Use Only

TU COMPLIANCE ASSESSMENT

| Pos. | Tested and checked items | TU clause No. |  | Compliance check results |
| --- | --- | --- | --- | --- |
|  |  | technical requirements | testing methods |  |
| 1 | Necessary documents availability | 1.1 | 5.6 | Compliant |
| 2 | Equipment quantity, marking and packaging | 1.3 1.4 1.5 | 5.6 | Compliant |
| 3 | Raw materials, materials and purchased items | 1.2 | 5.7 | Compliant |
| 4 | Size | 1.1.1 1.5.8 | 5.8 | Compliant |
| 5 | Weight | 1.1.1 1.5.9 | 5.9 | Compliant |
| 6 | AC powered operation | 1.1.2 | 5.10 | Compliant |
| 7 | Maximum nominal input | 1.1.3 | 5.11 | Compliant |
| 8 | Design proof | 1.1.4 1.1.5 1.1.18 | 5.12 | Compliant |
| 9 | Thermal block and heated lid temperature | 1.1.5 | 5.13 | Compliant |
| 10 | Software-controlled operation | 1.1.6 | 5.14 | Compliant |
| 11 | Time to set the operating mode | 1.1.7 | 5.15 | Compliant |
| 12 | Nonstop continuous run time | 1.1.7 | 5.15 | Compliant |
| 13 | Exterior inspection | 1.1.8 | 5.16 | Compliant |
| 20 | Assembly | 1.1.16 | 5.24 | Compliant |
| 23 | Electromagnetic compatibility | 2.3 | 5.27 | Compliant |
| 24 | Acoustic output | 2.4 | 5.28 | Compliant |

TRANSPORTATION AND STORAGE

The Amplifier is transported by all types of covered transport in accordance with the rules of cargo transportation applicable to such a type of transport.  Vehicles must be covered, dry and clean.

The Amplifier is transported in manufacturer’s packaging at an ambient temperature of -40°C to +65°C and relative humidity of 10% to 90% (non-condensing).

During loading, unloading, transportation of and storage, measures must be taken to protect the Amplifier from mechanical damage and contamination.

The Amplifier in manufacturer’s packaging must be stored in well-ventilated rooms at temperatures from -20°C to + 30°C and relative humidity from 30% to 85% (non-condensing).  Atmospheric pressure must be from 53.0 to 106.0 kPa.

DISPOSAL

The Amplifier that has become unusable, including due to the expiration of its shelf life, must be disposed of.

The Amplifier must be disposed of subject to SanPiN 2.1.7.2790 (applicable to hazard class “A” of health-care waste).

WARRANTY

The manufacturer warrants that the Amplifier meets the requirements of TU 26.51.53-003-06931260-2020 provided that the operating (usage), transportation and storage conditions are observed.

The warranty period is 12 (twelve) months from the initial start-up.

The service life of the Amplifier is 5 (five) years if used in full compliance with the Handbook.

The warranty storage life of the Amplifier is 3 (three) years of the date of manufacture, provided that it is stored under the aforementioned conditions.

Scheduled maintenance of the Amplifier must be carried out by an authorized service provider.

Warranty claims should be addressed to the manufacturer:

EVOTECH-MIRAI GENOMICS LLC

Address: 3B Akademika Artsimovicha St, Room 11, Moscow 117437, Russia

e-mail: mail@evotech-mg.com

Statement on Compliance:

Isothermal Amplifier for SARS-CoV-2 nucleic acid detection according to TU 26.51.53-003-06931260-2020

Serial number_______________________

Date of production 	_______________________

complies with TU 26.51.53-003-06931260-2020.

Head of Biological and Technical Control Department, EVOTECH-MIRAI GENOMICS LLC/