APPROVED BY

Lenar Valeev

General Director

EVOTECH-MIRAI GENOMICS LLC

________________

__________________ 20____

DATASHEET No.____

of ____________

Pretreatment Device for the nucleic acids extraction from the biological samples for SARS-CoV-2 isothermal amplification performance 
according to TU 26.51.53-001-06931260-2020.

Registration Certificate for a Medical Device

No. РЗН (RZN) 2020/10089 dated April 17, 2020

2020

MANUFACTURER

Manufacturer:

EVOTECH-MIRAI GENOMICS Limited Liability Company

(EVOTECH-MIRAI GENOMICS LLC)

Address: 3B Akademika Artsimovicha St, Room 11, Moscow 117437, Russia

The production site is located at:

Joint-Stock Company «Production Association «Zavod imeni Sergo» (POZIS) JSC POZIS, 4, Privokzalnaya Street, Zelenodolsk, Republic of Tatarstan, Russia 422546

LIST OF EQUIPMENT

| Pos. | Item | Quantity |
| --- | --- | --- |
| 1.1 | Pretreatment Device for the nucleic acids extraction from the biological samples for SARS-CoV-2 isothermal amplification performance according to TU 26.60.12-001-06931260-2020 (the “Pretreatment Device”) | 1 |
| 1.2 | Interface unit | 1 |
| 1.3 | Tray | 1 |
| 1.4 | Rack 1 (plugged unit) | 1 |
| 1.5 | Rack 2 (filtration unit) | 1 |
| 1.6 | Rack 3 (eluate collection unit) | 1 |
| 1.7 | Waste collection container with a lid | 1 |
| 1.8 | Bacterial air filter | 2 |
| 1.9 | Connection tube 1.8 m | 1 |
| 1.10 | Connection tube 0.3 m | 2 |
| 1.11 | Connecting wire | 1 |
| 1.12 | Compressor | 1 |
| 1.13 | Compressor connecting wire | 1 |
| 1.14 | Power cord | 1 |
| 1.15 | Operational documents: |  |
| 1.15.1 | Handbook | 1 |
| 1.15.2 | Datasheet | 1 |

SPECIFICATIONS

Weight (Pretreatment Device and equipment) - 1.4±0.1 kg.

Size:

– Length, mm: 210 ±10%

– Width, mm: 113 ±10%

– Height, mm: 174 ±10%.

The Pretreatment Device operates on AC 220±23V at the frequency of 50Hz.  Power supply unit specifications: input voltage is АС 100-240V, output voltage is DC 24V at 4.6А

Maximum power consumption: 24W.

Color-coded stage indicators:

- Lysis (SSB) - yellow indicator

- Washing 1 (WS1) - yellow indicator

- Washing 2 (WS2) - yellow indicator

- Drying (DRYING) - blue indicator

- Elution (ELUTION) - red indicator.

Procedure stages - time:

- Lysis (SSB) - 20 seconds

- Washing 1 (WS1) - 20 seconds

- Washing 2 (WS2) - 20 seconds

- Drying (DRYING) - 40 seconds

- Elution (ELUTION) - 20 seconds.

Tolerance ±10%

Negative pressure (vacuum):

from -10 to -30 kPa.

The maximum allowable time to establish the operating mode of the Pretreatment Device after switching on is no more than 60 seconds.

The nonstop continuous run time is minimum 8 hours.

Catalogue number

A5465 - Medical Device

А5465RUO - Research Use Only

TU COMPLIANCE ASSESSMENT

| Pos. | Tested and checked items | TU clause No. |  | Compliance check results |
| --- | --- | --- | --- | --- |
|  |  | technical requirements | testing methods |  |
|  | Necessary documents availability | 1.1 | 5.6 | Compliant |
|  | Equipment quantity, marking and packaging | 1.3 1.4 1.5 | 5.6 | Compliant |
|  | Raw materials, materials and purchased items | 1.2 | 5.7 | Compliant |
|  | Size | 1.1.1 | 5.8 | Compliant |
|  | Weight | 1.1.1 | 5.9 | Compliant |
|  | Operation at voltage change | 1.1.2 | 5.10 | Compliant |
|  | Power consumption | 1.1.3 | 5.11 | Compliant |
|  | Color-coded indicators | 1.1.4 | 5.12 | Compliant |
|  | Procedure stages - running time | 1.1.5 | 5.13 | Compliant |
|  | Negative pressure (vacuum) | 1.1.6 | 5.14 | Compliant |
|  | Time to set the operating mode | 1.1.7 | 5.15 | Compliant |
|  | Nonstop continuous run time | 1.1.7 | 5.15 | Compliant |
|  | Exterior inspection | 1.1.8 | 5.16 | Compliant |
|  | Assembly | 1.1.16 | 5.24 | Compliant |
|  | Electromagnetic compatibility | 2.3 | 5.27 | Compliant |
|  | Acoustic output | 2.4 | 5.28 | Compliant |

TRANSPORTATION AND STORAGE

The Pretreatment Device is transported by all types of covered transport in accordance with the rules of cargo transportation applicable to such a type of transport. Vehicles must be covered, dry and clean.

The Pretreatment Device is transported in manufacturer’s packaging at -40°C to +65°C and relative humidity of 10% to 90% (non-condensing).

During loading, unloading, transportation of and storage, measures must be taken to protect the Pretreatment Device from mechanical damage and contamination.

The Pretreatment Device in manufacturer’s packaging must be stored in well-ventilated rooms at                   -20°C to + 40°C and relative humidity of 30% to 85% (non-condensing).  Atmospheric pressure must be from 53.0 to 106.0 kPa.

DISPOSAL

The Pretreatment Device that has become unusable, including due to the expiration of its shelf life, must be disposed of.

The Pretreatment Device must be disposed of subject to SanPiN 2.1.7.2790 (applicable to hazard class “A” of health-care waste).

WARRANTY

The manufacturer warrants that Pretreatment Device meets the requirements of TU 26.51.53-001-06931260-2020 provided that the operating (usage), transportation and storage conditions are observed.

The warranty period is 12 (twelve) months from the initial start-up.

The service life of the Pretreatment Device is 5 (five) years if used in full compliance with the Handbook.

The warranty storage life of the Pretreatment Device is 3 (three) years of the date of manufacture, provided that it is stored under the aforementioned conditions.

Scheduled maintenance of the Pretreatment Device must be carried out by an authorized service provider.

Warranty claims should be addressed to the manufacturer: EVOTECH-MIRAI GENOMICS LLC at

3B Akademika Artsimovicha St, Room 11, Moscow 117437, Russia

Statement on Compliance:

The Pretreatment Device for the nucleic acids extraction from the biological samples for SARS-CoV-2 isothermal amplification performance according to TU 26.60.12-001-06931260-2020 by EVOTECH-MIRAI GENOMICS LLC

Serial number	_______________________

Date of production 	_______________________

complies with TU 26.51.53-001-06931260-2020.

Head of Biological and Technical Control Department, EVOTECH-MIRAI GENOMICS LLC	/