APPROVED BY

Lenar Valeev

General Director

EVOTECH-MIRAI GENOMICS LLC

________________

____________________ 2020

DATASHEET No.____

of ____________

Reagent kit for the SARS-CoV-2 RNA nucleic acids extraction from biological material in different versions according to TU 21.10.60-002-06931260-2020

Registration Certificate for a Medical Device

No.2020/10087 of April 17, 2020

2020

| ## Reagent kit for the SARS-CoV-2 RNA nucleic acids extraction from biological material in different versions according to TU 21.10.60-002-06931260-2020 (the “Kit”) ## Version II. 96 detections. |  |
| --- | --- |
| ## Batch No. _______ | ## Date of production _______________________ |
| ## Number of Kits per batch ____ |  |

MANUFACTURER

EVOTECH-MIRAI GENOMICS Limited Liability Company

(EVOTECH-MIRAI GENOMICS LLC)

Address: 3B Akademika Artsimovicha St, Room 11, Moscow 117437, Russia

The production site is located at:

Sigma Lab LLC

Address: 42, Bolshoy Boulevard, Bldg. 1, Floor 3, Room 785, Territory of the Skolkovo Innovation Center, Moscow, 143026 Russia

LIST OF EQUIPMENT

| Pos. | Item | Quantity, pcs |
| --- | --- | --- |
| 1. | Version II. 96 detections: | --- |
| 1.1. | SSB solution, transparent colorless liquid, 50 ml | 2 |
| 1.2. | WS1 solution, transparent colorless liquid, 50 ml | 2 |
| 1.3. | WS2 solution, transparent colorless liquid, 50 ml | 6 |
| 1.4. | EB solution, transparent colorless liquid, 5 ml | 2 |
| 1.5. | Operational documents | --- |
| 1.5.1 | Handbook | 1 |
| 1.5.2 | Datasheet | 1 |

SPECIFICATIONS

The Kit components are disposable.  The Kit does not require maintenance and calibration.

SSB solution is a transparent colorless liquid, which comes in 50ml plastic tubes with screw caps.

WS1 solution is a transparent colorless liquid, which comes in 50ml plastic tubes with screw caps.

WS2, solution is a transparent colorless liquid, which comes in 50ml plastic tubes with screw caps.

EB solution is a transparent colorless liquid, which comes in 5ml plastic tubes with screw caps.

Contents:

| Name of Kit components | Component | Final concentration in the reagent | Compliance assessment |
| --- | --- | --- | --- |
| SSB solution | Guanidine thiocyanate | 2.4 mol | Compliant |
|  | Sodium polyacrylate | 0.0025% | Compliant |
|  | Buffer Tris-HCl, рН 7.4 | 25 mmol | Compliant |
|  | Triton X-100 | 0.5% | Compliant |
|  | Ethanol | 33% | Compliant |
| WS1 solution | Guanidine hydrochloride | 2.4 mol | Compliant |
|  | Ethanol | 50% | Compliant |
| WS2 solution | Ethanol | 95% | Compliant |
| EB solution | Nuclease-free distilled water | 100% | Compliant |

Number of detections

The Kit is designed for 96 detections.

Microbial limits – 2.84

FUNCTIONAL SPECIFICATIONS

The Kit, used in conjugation with the Pretreatment Device, yields elute solution containing purified RNA. Based on the ratio of absorbed 260 nm/ 280nm wavelength the purity of solution is within the range of 2.7 - 2.9. Microbial limit - 2.84.

Known interference:

Substances in concentrations as shown in the table below do not affect RNA extraction efficiency from biological samples:

| Interferent | Concentration in a sample |
| --- | --- |
| Mucin* | 5% v/v |

*	Mucin from porcine stomach Type II, Sigma, M2378-100G, Lot # SLCC7713 is used.

Reagent consumption:

| Reagent | Function | Required volume per sample |
| --- | --- | --- |
| SSB solution | Lysis buffer | 1.0ml |
| WS1 solution | Washing solution | 1.0ml |
| WS2 solution | Washing solution | 1.0ml |
| EB solution | Eluting solution | 100μl |

TU COMPLIANCE ASSESSMENT

| Tested and checked items | TU clause No. |  | Compliance check results |
| --- | --- | --- | --- |
|  | technical requirements | testing methods |  |
| Equipment documentation, materials, quaintity, marking and packaging | 1.1.1 1.2 1.3 1.4 1.5 | 5.2.1 | Compliant |
| Size | 1.1.2 | 5.2.2 | Compliant |
| Tube liquid volume | 1.1.2 | 5.2.2 | Compliant |
| Tube visual inspection, content inspection | 1.1.2 1.1.5 | 5.2.3 | Compliant |
| Leakage | 1.1.5 | 5.2.4 | Compliant |
| Composition | 1.1.6 | 5.2.5 | Compliant |
| Extraction quality | 1.1.7 | 5.2.6 | Compliant |

TRANSPORTATION AND STORAGE

The Kit is transported by all types of covered transport in accordance with the rules of cargo transportation applicable to such a type of transport, at temperatures from +2°С to +30°С, in an environment free from hostile conditions, direct sunlight and moisture.  Freezing of the Kit is not allowed.

During loading, unloading, transportation of and storage, measures must be taken to protect the Kit container from mechanical damage, exposure to atmospheric precipitation and aggressive environments.

The Kit in manufacturer’s packaging must be stored at temperatures from +2°C to +30°C under conditions excluding the effect of aggressive environments, direct sunlight and moisture.  Freezing of the Kit is not allowed.

DISPOSAL

Used components, which have contaced biological samples, are subject to desinfection and then must be stored in waste collection containers or plastic bags and disposed of as Class C extremely hazardous waste according to SanPiN 2.1.7.2790-10.  Kits that have become unusable, inclusive of expired Kits and Kits that were opened but not used, are disposed of as Class D according to SanPiN 2.1.7.2790-10 and MU 1.3.2569-09. The packaging is disposed of as Class A according to SanPiN 2.1.7.2790-10.

WARRANTY

The manufacturer warrants: the Kit functional specifications meet the requirements of specifications and operational documents within the established shelf life (12 months), provided that the transportation and storage conditions are observed.

Warranty claims should be addressed to the manufacturer:

EVOTECH-MIRAI GENOMICS LLC

Address: 3B Akademika Artsimovicha St, Room 11, Moscow 117437, Russia

e-mail: mail@evotech-mg.com.

Any side effects not listed in the Kit Instruction for Use, undesirable reactions, facts or conditions posing a threat to the health and life of people and medical staff and associated with the use of the Kit should be reported to the manufacturer, EVOTECH-MIRAI GENOMICS LLC, at the above address, as well as to the relevant governmental supervising authority as required by the applicable law.

Statement on Compliance:

Reagent kit for the SARS-CoV-2 RNA nucleic acids extraction from biological material in different versions according to TU 21.10.60-002-06931260-2020.

Version II. 96 detections

Batch No.	LOT

Date of production 	_______________________

complies with TU 21.10.60-002-06931260-2020.

Head of Biological and Technical Control Department, EVOTECH-MIRAI GENOMICS LLC	/